## Sheet1
Attribute charts
Chart | Uses these dataset columns | Notes
p | Batch ID, Timestamp, n Inspected, Defectives (np) (or Defect Rate (p)) | Uses varying n per batch OK
np | Batch ID, Timestamp, n Inspected, Defectives (np) | Works best when n ~ constant; dataset can still plot
c | Batch ID, Timestamp, Defects (c) | Assumes constant opportunity window
u | Batch ID, Timestamp, Defects (c), Opportunities (or Defects per Unit (u)) | Preferred if opportunities vary
Variable charts
Chart | Uses these dataset columns | Notes
X̄–R | Subgroup ID, Timestamp, RTD Temperature (C) or pH Level or Absorbance (OD) | Subgroup size is Sample No count (you have 5)
X̄–S | same as above | Best when subgroup size ≥ 10 (still works)
I–MR | Timestamp + one of the measurements above (or residuals) | Uses individual points
1) RTD Temperature (critical process parameter)
Manufacturing context | Why this is best practice | Recommended chart | Dataset variable(s)
Real-time sensor stream | Detects sudden shifts, drift, heater failures, sensor noise | I–MR | RTD Temperature (C)
Batch with 3–5 samples | Gold standard for short subgroups | X̄–R | RTD Temperature (C)
Batch with ≥10 samples | Better σ estimate for larger subgroups | X̄–S | RTD Temperature (C)
Digital twin validation | Flags model deviation → smart-manufacturing anomaly | I–MR | Temp Residual
Stability trending | Required for thermal process validation evidence | I–MR (baseline limits) | RTD Temperature (C)
2) pH (chemistry-sensitive parameter)
Manufacturing context | Why this is best practice | Recommended chart | Dataset variable(s)
Inline / frequent pH checks | Detects probe drift, calibration loss | I–MR | pH Level
Subgrouped pH samples | Identifies reagent prep instability | X̄–R | pH Level
Digital twin comparison | Confirms chemistry model accuracy | I–MR | pH Residual
Lot-to-lot assessment | Better sensitivity for formulation shifts | X̄–S | pH Level
⚠️ Regulatory note: pH often has tight specs → spec lines (red) are critical, but control limits must still exist.
3) Optical Density (OD) / Absorbance (assay performance)
Manufacturing context | Why this is best practice | Recommended chart | Dataset variable(s)
Per-run OD readings | Detects wash failures, reader drift | I–MR | Absorbance (OD)
Plate or batch averages | Separates within-plate vs between-plate variation | X̄–R | Absorbance (OD)
Model-based monitoring | Digital twin anomaly detection | I–MR | OD Residual
Release trending | Required for assay consistency evidence | X̄–S | Absorbance (OD)
📌 FDA insight: OD instability often correlates with false negatives/positives → high patient risk (ISO 14971).
4) Defectives (pass/fail outcomes)
Data situation | Why | Recommended chart | Dataset variable(s)
Variable sample size | Most common FDA-reviewed attribute chart | p chart | Defect Rate (p) + n Inspected
Constant sample size | Easier interpretation for operators | np chart | Defectives (np) + n Inspected
Incoming/release inspection | Strong link to CAPA initiation | p chart | Defect Rate (p)
⚠️ Never use p/np charts for sensor data — auditors will flag this.
5) Defects (multiple defects per unit)
Data situation | Why | Recommended chart | Dataset variable(s)
Fixed inspection area | Classic Poisson assumption | c chart | Defects (c)
Variable opportunities | Required when opportunity count changes | u chart | Defects (c) / Opportunities
Contamination monitoring | Sensitive to hygiene/process breakdowns | u chart | Defects per Unit (u)
6) Smart Manufacturing / Quality 5.0 view (recommended defaults)
Objective |  | Chart | Variable
Process stability |  | X̄–R | RTD Temperature (C)
Chemistry control |  | X̄–R | pH Level
Assay performance |  | X̄–R | Absorbance (OD)
Real-time alarms |  | I–MR | Sensor values
Model trustworthiness |  | I–MR | Residuals
Product quality |  | p / u | Defectives / defects
7) What auditors like to see (practical guidance)
Variables data first (X̄–R, I–MR) → process control
Attributes second (p/u) → product outcome
Control limits ≠ Specification limits (you’ve already enforced this visually ✔)
Residual charts → strong evidence of Industry 4.0 / 5.0 maturity